## Sheet0
id | collection | dc.contributor.author | dc.date.accessioned | dc.date.available | dc.date.issued | dc.description.abstract[en] | dc.identifier.uri | dc.language | dc.rights.uri | dc.subject.ddc | dc.subject[en] | dc.title | dc.type | dcterms.accessRights.openaire | dcterms.bibliographicCitation.articlenumber | dcterms.bibliographicCitation.doi | dcterms.bibliographicCitation.journaltitle | dcterms.bibliographicCitation.originalpublishername | dcterms.bibliographicCitation.pmid | dcterms.bibliographicCitation.volume | dcterms.isPartOf.eissn | refubium.affiliation | refubium.resourceType.isindependentpub
28159d01-2252-4ccc-b6fc-0acb90a99e49 | fub188/15 | Rogasch, Julian M.¦¦Suleiman, Said¦¦Hofheinz, Frank¦¦Bluemel, Stephanie¦¦Lukas, Mathias¦¦Amthauer, Holger¦¦Furth, Christian | 2020-02-07T10:34:10Z | 2020-02-07T10:34:10Z | 2020 | BACKGROUND: Bayesian penalized likelihood reconstruction for PET (e.g., GE Q.Clear) aims at improving convergence of lesion activity while ensuring sufficient signal-to-noise ratio (SNR). This study evaluated reconstructed spatial resolution, maximum/peak contrast recovery (CRmax/CRpeak) and SNR of Q.Clear compared to time-of-flight (TOF) OSEM with and without point spread function (PSF) modeling. METHODS: The NEMA IEC Body phantom was scanned five times (3 min scan duration, 30 min between scans, background, 1.5-3.9 kBq/ml F18) with a GE Discovery MI PET/CT (3-ring detector) with spheres filled with 8-, 4-, or 2-fold the background activity concentration (SBR 8:1, 4:1, 2:1). Reconstruction included Q.Clear (beta, 150/300/450), "PSF+TOF4/16" (iterations, 4; subsets, 16; in-plane filter, 2.0 mm), "OSEM+TOF4/16" (identical parameters), "PSF+TOF2/17" (2 it, 17 ss, 2.0 mm filter), "OSEM+TOF2/17" (identical), "PSF+TOF4/8" (4 it, 8 ss, 6.4 mm), and "OSEM+TOF2/8" (2 it, 8 ss, 6.4 mm). Spatial resolution was derived from 3D sphere activity profiles. RC as (sphere activity concentration [AC]/true AC). SNR as (background mean AC/background AC standard deviation). RESULTS: Spatial resolution of Q.Clear150 was significantly better than all conventional algorithms at SBR 8:1 and 4:1 (Wilcoxon, each p < 0.05). At SBR 4:1 and 2:1, the spatial resolution of Q.Clear300/450 was similar or inferior to PSF+TOF4/16 and OSEM+TOF4/16. Small sphere CRpeak generally underestimated true AC, and it was similar for Q.Clear150/300/450 as with PSF+TOF4/16 or PSF+TOF2/17 (i.e., relative differences < 10%). Q.Clear provided similar or higher CRpeak as OSEM+TOF4/16 and OSEM+TOF2/17 resulting in a consistently better tradeoff between CRpeak and SNR with Q.Clear. Compared to PSF+TOF4/8/OSEM+TOF2/8, Q.Clear150/300/450 showed lower SNR but higher CRpeak. CONCLUSIONS: Q.Clear consistently improved reconstructed spatial resolution at high and medium SBR compared to PSF+TOF and OSEM+TOF, but only with beta = 150. However, this is at the cost of inferior SNR with Q.Clear150 compared to Q.Clear300/450 and PSF+TOF4/16/PSF+TOF2/17 while CRpeak for the small spheres did not improve considerably. This suggests that Q.Clear300/450 may be advantageous for the 3-ring detector configuration because the tradeoff between CR and SNR with Q.Clear300/450 was superior to PSF+TOF4/16, OSEM+TOF4/16, and OSEM+TOF2/17. However, it requires validation by systematic evaluation in patients at different activity and acquisition protocols. | https://refubium.fu-berlin.de/handle/fub188/26608¦¦http://dx.doi.org/10.17169/refubium-26365 | eng | https://creativecommons.org/licenses/by/4.0/ | 600 Technik, Medizin, angewandte Wissenschaften::610 Medizin und Gesundheit::610 Medizin und Gesundheit | PET¦¦Image reconstruction¦¦Spatial resolution¦¦Contrast recovery¦¦Signal-to-noise ratio,¦¦TOF¦¦PSF¦¦Q.Clear¦¦GE Discovery MI | Reconstructed spatial resolution and contrast recovery with Bayesian penalized likelihood reconstruction (Q.Clear) for FDG-PET compared to time-of-flight (TOF) with point spread function (PSF) | Wissenschaftlicher Artikel | open access | 2 | 10.1186/s40658-020-0270-y | EJNMMI Physics | SpringerOpen | 31925574 | 7 | 2197-7364 | Charité - Universitätsmedizin Berlin | no