## Sheet0
id | collection | dc.contributor.author | dc.date.accessioned | dc.date.available | dc.date.issued | dc.description.abstract[en] | dc.format.extent | dc.identifier.uri | dc.language | dc.rights.uri | dc.subject.ddc | dc.subject[en] | dc.title | dc.type | dcterms.accessRights.openaire | dcterms.bibliographicCitation.doi | dcterms.bibliographicCitation.journaltitle | dcterms.bibliographicCitation.number | dcterms.bibliographicCitation.pageend | dcterms.bibliographicCitation.pagestart | dcterms.bibliographicCitation.url | dcterms.bibliographicCitation.volume | dcterms.isPartOf.eissn | refubium.affiliation | refubium.affiliation.other | refubium.resourceType.isindependentpub | refubium.resourceType.provider
5a6c75f4-26c8-4f3a-9e48-0b13fdee8e48 | fub188/16 | Abad, Carlos¦¦Jegielka, Dennis¦¦Aloysius, Allen¦¦Recknagel, Sebastian | 2025-10-06T11:20:21Z | 2025-10-06T11:20:21Z | 2025 | Accurate nitrate and nitrite data support water-quality regulation, yet routine methods rely on external calibration and rarely achieve SI traceability. We report a calibration-free determination of nitrate and nitrite by combining isotope dilution with high-resolution continuum-source graphite furnace molecular absorption spectrometry (ID-HR-CS-GF-MAS). A 15N-enriched nitrate spike (its concentration verified by reverse isotope dilution against the standard reference material NIST 3185) provides the SI link, and it is gravimetrically added to samples; nitrate and residual nitrite are converted in situ to nitric oxide (NO), whose 215 nm band is recorded at a pixel resolution of λ/Δλ ≈ 140 000. The 0.2127 nm shift between 14NO and 15NO electronic spectra is resolved, and a three-latent-variable partial least squares regression model yields the 15N/14N ratio with 0.3% precision. Instrumental LoD values of 4.8 ng (14N) and 3.2 ng (15N) translate to a method LoD of 4.8 ng of nitrogen (equivalent to 1.05 mg L−1 NO3− for a 20 μL aliquot). The furnace program allows for successive drying/pyrolysis loops, so additional 20 μL aliquots can be layered onto the graphite platform. Alternatively, a 10 mL anion-exchange solid-phase extraction step concentrates nitrate and nitrite fivefold, allowing for the analysis of even lower sample concentrations. Results for four certified reference materials (2.9 to 1000 mg L−1 NO3−) agreed with certified values, giving relative expanded uncertainties of 2 to 4%. Analysis of twenty Berlin surface-water samples revealed concentrations ranging from 0.10 to 7.3 mg L−1 NO3−, indicating that the Panke River and Teltow Canal are the primary sources of nitrogen. ID-HR-CS-GF-MAS thus delivers ID-MS-level accuracy in a few minutes per run with bench-top optics, and, with optional on-platform or SPE pre-concentration, extends SI-traceable nitrate/nitrite monitoring into the low-ng regime. | 10 Seiten | https://refubium.fu-berlin.de/handle/fub188/49489¦¦http://dx.doi.org/10.17169/refubium-49211 | eng | https://creativecommons.org/licenses/by/4.0/ | 500 Naturwissenschaften und Mathematik::540 Chemie::540 Chemie und zugeordnete Wissenschaften | analysis¦¦nitrate¦¦nitrite | SI-traceable total analysis of nitrate and nitrite by isotope dilution optical spectroscopy and its application to Berlin surface waters | Wissenschaftlicher Artikel | open access | 10.1039/D5JA00252D | Journal of Analytical Atomic Spectrometry | 10 | 2701 | 2692 | https://doi.org/10.1039/D5JA00252D | 40 | 1364-5544 | Biologie, Chemie, Pharmazie | Institut für Chemie und Biochemie:::bdaf1c47-f341-464f-b012-0d24067d34b8:::600 | no | WoS-Alert