## Sheet0
id | collection | dc.contributor.author | dc.date.accessioned | dc.date.available | dc.date.issued | dc.description.abstract[en] | dc.identifier.uri | dc.language | dc.rights.uri | dc.subject.ddc | dc.subject[en] | dc.title | dc.type | dcterms.accessRights.openaire | dcterms.bibliographicCitation.doi | dcterms.bibliographicCitation.journaltitle | dcterms.bibliographicCitation.number | dcterms.bibliographicCitation.originalpublishername | dcterms.bibliographicCitation.pageend | dcterms.bibliographicCitation.pagestart | dcterms.bibliographicCitation.pmid | dcterms.bibliographicCitation.volume | dcterms.isPartOf.eissn | dcterms.isPartOf.issn | refubium.affiliation | refubium.funding | refubium.resourceType.isindependentpub
cd4e3df4-d33d-4614-8495-d48c65641ed1 | fub188/15 | Mähler, Anja¦¦Schütte, Till¦¦Steiniger, Jochen¦¦Boschmann, Michael | 2025-11-18T12:20:11Z | 2025-11-18T12:20:11Z | 2023 | Purpose We present a methodological overview of a respiration chamber at the Experimental and Clinical Research Center in Berlin, Germany. Since 2010, we investigated 750 healthy subjects and patients with various diseases. We routinely measure resting energy expenditure (REE), dietary-induced thermogenesis, and activity energy expenditure.Methods The chamber is a pull calorimeter with a total volume of 11,000 L. The majority of measurements is done with a flow rate of 120 L/min, yielding a favorable time constant of 1.53 h. The gas analysis system consists of two paramagnetic O-2 sensors and two infrared CO2 sensors, one for incoming and one for outgoing air samples. O-2 and CO2 sensors are calibrated simultaneously before each measurement with a 6 min calibration routine. To verify the accuracy of the whole the calorimetric system, it is validated every 2 weeks by 2 h acetone burning tests.Results Validation factors (calculated/measured) of 20 representative 2 h acetone burning tests were 1.03 +/- 0.03 for ?VO2, 1.02 +/- 0.02 for ?VCO2, 0.99 +/- 0.02 for RER, and 1.03 +/- 0.03 for EE. Four repeated 60 min REE measurements of a healthy woman showed variabilities of 231.9 +/- 4.8 ml/min for ?VO2 (CV 2.1%), 166.0 +/- 6.3 ml/min for ?VCO2 (CV 3.8%), 0.73 +/- 0.03 for RER (CV 4.6%), and 4.55 +/- 0.07 kJ/min for EE (CV 1.6%)..Conclusions The data presented show that our respiration chamber produces precise and valid EE measurements with an exceptionally fast responsiveness. | https://refubium.fu-berlin.de/handle/fub188/50411¦¦http://dx.doi.org/10.17169/refubium-50137 | eng | https://creativecommons.org/licenses/by/4.0/ | 600 Technik, Medizin, angewandte Wissenschaften::610 Medizin und Gesundheit::610 Medizin und Gesundheit | Whole-room calorimeter¦¦Respiration chamber¦¦Energy expenditure¦¦Diet-induced thermogenesis | The Berlin-Buch respiration chamber for energy expenditure measurements | Wissenschaftlicher Artikel | open access | 10.1007/s00421-023-05164-w | European Journal of Applied Physiology | 6 | Springer Nature | 1368 | 1359 | 36849666 | 123 | 1439-6327 | 1439-6319 | Charité - Universitätsmedizin Berlin | Springer Nature DEAL | no