## Sheet0
id | collection | dc.contributor.author | dc.date.accessioned | dc.date.available | dc.date.issued | dc.description.abstract[en] | dc.format.extent | dc.identifier.sepid | dc.identifier.uri | dc.language | dc.rights.uri | dc.subject.ddc | dc.subject[en] | dc.title | dc.type | dcterms.accessRights.openaire | dcterms.bibliographicCitation.doi | dcterms.bibliographicCitation.journaltitle | dcterms.bibliographicCitation.number | dcterms.bibliographicCitation.originalpublishername | dcterms.bibliographicCitation.pageend | dcterms.bibliographicCitation.pagestart | dcterms.bibliographicCitation.url | dcterms.bibliographicCitation.volume | dcterms.isPartOf.eissn | refubium.affiliation | refubium.affiliation.other | refubium.funding | refubium.note.author[de] | refubium.resourceType.isindependentpub
b5a04c43-a987-4962-a9f6-478fdfa70ee3 | fub188/16 | Venjakob, Peter L.¦¦Bretzinger, Lukas¦¦Borchardt, Stefan¦¦Weber, Corinna¦¦Heuwieser, Wolfgang | 2023-08-10T10:13:58Z | 2023-08-10T10:13:58Z | 2023 | The objective of this study was to compare measurements of nonesterified fatty acids (NEFA) between the gold standard diagnostic laboratory method and a handheld NEFA meter (Qucare Pro meter, DFI Co. Ltd.). Three experiments were conducted to study the usability of the meter. In experiment 1 we compared results of the meter obtained from measurements in serum and whole blood with results of the gold standard method. Based on the results of experiment 1 we compared the results measured by the meter in whole blood with results obtained from the gold standard method on a larger scale, as we wanted to omit the step of centrifugation with the cow-side test. In experiment 3 we determined the influence of ambient temperature on measurements. Overall, blood samples of 231 cows were collected between 14 and 20 d in milk. The Spearman correlation coefficients (ρ) were calculated and Bland-Altman plots were created to compare the accuracy of the NEFA meter with the gold standard. In addition, in experiment 2 receiver operating characteristic (ROC) curve analyses were performed, to define thresholds for the NEFA meter to detect cows with a NEFA concentration above 0.3, 0.4, and 0.7 mEq/L. In experiment 1, there was a high correlation between NEFA concentrations in whole blood and serum determined by the NEFA meter and the gold standard (ρ = 0.90 for measurement in whole blood; ρ = 0.93 for measurement in serum). In experiment 2, the measurement in whole blood with the NEFA meter was compared with the gold standard. Despite a lower correlation (ρ = 0.79) the ROC curve analyses revealed a high specificity and a moderate sensitivity for lower cut-points (i.e., 0.3 and 0.4 mEq/L). The NEFA meter underestimated especially high concentration of >0.7 mEq/L. Considering thresholds of 0.3, 0.4, and 0.7 mEq/L measured by the gold standard test, sensitivity and specificity were 59.1% and 96.7%, 79.0% and 95.4%, and 86.4% and 95.6%, respectively, when using 0.3, 0.3, and 0.4 mEq/L as thresholds for the NEFA meter. Accuracy was 74.1%, 88.3%, and 93.8% for the 3 thresholds tested. Experiment 3 showed that measurements should be conducted at approximately 21°C (ρ = 0.73) as correlations were poor at 6.2°C and 15.1°C (ρ = 0.18 and 0.22, respectively). | 7 Seiten | 94156 | https://refubium.fu-berlin.de/handle/fub188/40420¦¦http://dx.doi.org/10.17169/refubium-40141 | eng | https://creativecommons.org/licenses/by/4.0/ | 600 Technik, Medizin, angewandte Wissenschaften::630 Landwirtschaft::630 Landwirtschaft und verwandte Bereiche | nonesterified fatty acids¦¦NEFA¦¦handheld NEFA meter | Evaluation of the diagnostic accuracy of a point-of-care device to measure concentrations of nonesterified fatty acids in serum and whole blood | Wissenschaftlicher Artikel | open access | 10.3168/jdsc.2022-0269 | JDS Communications | 3 | Elsevier | 224 | 219 | https://doi.org/10.3168/jdsc.2022-0269 | 4 | 2666-9102 | Veterinärmedizin | Tierklinik für Fortpflanzung:::ada5e16f-9f1e-4f3b-9bf6-01a242c43dd3:::600 | Publikationsfonds FU | Die Publikation wurde aus Open Access Publikationsgeldern der Freien Universität Berlin gefördert. | no